## Tests to be run
Name | Description | Method | Repetitions | Sample rate | Save flow profile | Save logged flow | Save logged volume | Save video | Write down measurements | Save logged pressure
Test 1 | Verification of flow by comparing actual position with logged position. | Measured by filming with high speed camera | Each case is repeated 3 times | Sample rate is 0.040s | x | x | x | x
1.1 | Sine wave, 1 period, frequency 1, amplitude 2
1.2 | Sine wave, 1 period, frequency 3, amplitude 20
1.3 | Sine wave, 1 period, frequency 6, amplitude 200
Test 2 | Verification of measured position by running a flow profile and comparing actual end position with logged end position | Distance between the sled and the syringe is measured with a caliper | Each case is repeated 3 times | Sample rate is 0.040s | x |  | x |  | x
2.1 | Square wave, 0.5 periods, frequency 6, amplitude 200, padded with zeros
2.2 | Square wave, 1 period, frequency 3, amplitude 20, padded with zeros
2.3 | Square wave, 1,5 periods, frequency 3, amplitude 20, padded with zeros
Test 3 | Verification of max pressure by running profiles with different amplitude and measuring pressure | Measured with external pressure meter | Each case is repeated 1 times | Sample rate is 0.040s |  | x |  |  | x | x (if possible)
3.1 | Sine wave 3 periods, frequency 3, amplitude 20
3.2 | Sine wave 3 periods, frequency 6, amplitude 200
3.3 | Square wave 3 periods, frequency 6, amplitude 150
Test 4 | Test of how well the recorded flow follows the inputted flow profile | Recorded flow is measured in the software | Each case is repeated 1 time | Sample rate is 0.040s | x | x
4.1 | Sine wave 3 periods, frequency 1, amplitude 2
4.2 | Sine wave 3 periods, frequency 3, amplitude 20
4.3 | Sine wave 3 periods, frequency 6, amplitude 200
4.4 | Triangle wave 3 periods, frequency 3, amplitude 20
4.5 | Square wave 3 periods, frequency 3, amplitude 20
Extra | Amp 200 F 14 10 rep

## Blad3
Test 2 |  | Innan | Efter | Data [ml] | ml | % error | dV [ml] | Medel | Median | STD | Medel/Vol | Median/Vol | STD/Vol | Area | 9.0792
1 | 1 | 97.25 | 26.4 | 65.895 | 64.3262 | 2.439% | 1.5688 | 0.85769 | 0.52365 | 0.50317 | 1.319% | 0.7858% | 0.7925%
| 2 | 96.5 | 23.1 | 67.165 | 66.6413 | 0.7858% | 0.52365 | Front - something
| 3 | 98.4 | 26 | 66.214 | 65.7334 | 0.7311% | 0.48057
2 | 1 | 73.3 | 73.05 | 0.177 | 0.22698 | -22.02% | -0.04998 | 0.029145 | 0.019812 | 0.068733 | 83.86% | 14.55% | 124.8% | Ointressanta siffror?
| 2 | 73.05 | 72.9 | 0.156 | 0.13619 | 14.55% | 0.019812 | Two way - Almost nothing. Similar error affecting both directions
| 3 | 72.9 | 72.85 | 0.163 | 0.045396 | 259.1% | 0.1176
3 | 1 | 72.9 | 50.6 | 20.747 | 20.2466 | 2.471% | 0.50038 | 0.30287 | 0.2664 | 0.14863 | 1.484% | 1.301% | 0.743%
| 2 | 67.25 | 44.25 | 21.024 | 20.8822 | 0.6792% | 0.14183 | Three way - something |  |  | 1.401% | 1.043% | 0.7677%
| 3 | 90.75 | 68.2 | 20.74 | 20.4736 | 1.301% | 0.2664
2.1 | Square wave, 0.5 periods, frequency 6, amplitude 200, padded with zeros
2.2 | Square wave, 1 period, frequency 3, amplitude 20, padded with zeros
2.3 | Square wave, 1,5 periods, frequency 3, amplitude 20, padded with zeros

## Blad2
Test 2 |  | Innan | Efter | Data [ml] | ml | % error | dV [ml] |  |  |  | Area | 9.0792
1 | 1 | 97.25 | 26.4 | 65.895 | 64.3262 | 2.439% | 1.5688 | 0.85769 | 1.319%
| 2 | 96.5 | 23.1 | 67.165 | 66.6413 | 0.7858% | 0.52365
| 3 | 98.4 | 26 | 66.214 | 65.7334 | 0.7311% | 0.48057
2 | 1 | 73.3 | 73.05 | 0.177 | 0.22698 | -22.02% | -0.04998 | 0.029145 | 83.86% | 0.5 | Front
| 2 | 73.05 | 72.9 | 0.156 | 0.13619 | 14.55% | 0.019812 |  |  | 1.0 | Back
| 3 | 72.9 | 72.85 | 0.163 | 0.045396 | 259.1% | 0.1176 |  |  | 1.5 | Front
3 | 1 | 72.9 | 50.6 | 20.747 | 20.2466 | 2.471% | 0.50038 | 0.30287 | 1.484%
| 2 | 67.25 | 44.25 | 21.024 | 20.8822 | 0.6792% | 0.14183
| 3 | 90.75 | 68.2 | 20.74 | 20.4736 | 1.301% | 0.2664
1.043%
Test 3 |  | Max Vol
2 | Av Ants | 65,8/65,7/67,82 kvar med init botten
3 | Av Ants | 74/71/70 kvar med init botten
1 | Av Ants | 13,2/13,5/13,5 kvar med init nästan i botten
Extra | Amp 200 F 14 10 rep |  |  |  | 2.2282 | Hz
3.4 | Max flow nästan 200
Vol ca 22-26
Av Ants
Vid låga flöden och snabb sampling påverkar upplösningen på motorn och gör flödessignalen brusig

## Sheet1
Test what? | Test how? | Test inputs | Repeat test (times) | Data to save
Actual flow vs recorded flow
Recorded flow vs flow profile
Recorded pressure during different flows
Actual end position vs recorded end position
Types of input
Constant flow | Square during 0.5 or less periods
Changing flow | Sine, Triangle and Square
Types of changing flow | Frequency
Rapid | 5 rad/s
Medium | 3 rad/s
Slow | 1 rad/s
Types of magnitudes
High | 200 ml/s
Medium | 20 ml/s
Low | 2 ml/s

## Blad1
ID | Description | Motivation | Measurement method | Measured Unit | Repetitions | Required mean value | Required standard deviation | Measured mean value | Measured standard deviation | Measured Value 1 | Measured Value 2 | Measured Value 3 | Measured value 4 | Measured Value 5
1. Flow accuracy
1.1
1.2
1.3
1.4
2. Volume accuracy
2.1 | Check recorded volume after constant flow of 200 ml/s for 0.3s. | Test that a fast acceleration and deceleration doesn't affect the reading of the end volume . | Measure start and end position with a caliper and compare with the volume recorded in the software. | percent | 5 | 0 | 5
2.2 | Check recorded volume after constant flow of 5 ml/s for 4s. | The resolution of the servo might be a problem at slow flows. | Measure start and end position with a caliper and compare with the volume recorded in the software. | percent | 5 | 0 | 5
2.3 | Check recorded volume after sine with amplitude 100 ml/s and frequency 6 rad/s for 3 periods. | Tests the backlash of the system for rapid changes. | Measure start and end position with a caliper and compare with the volume recorded in the software. | percent | 5 | 0 | 5
2.4 | Check recorded volume after sine with amplitude 10 ml/s and frequency 1 rad/s for 3 periods. | Tests the backlash of the system for slow changes. | Measure start and end position with a caliper and compare with the volume recorded in the software. | percent | 5 | 0 | 5
2.1 | Check outputted volume after constant flow of 200 ml/s for 0.3s. | Test how well the outputted volume corresponds to the theoretical volume | Measure start and end position with a caliper and compare with the volume recorded in the software. | ml | 5 | 60 | 3
2.2 | Check outputted volume after constant flow of 5 ml/s for 4s. | The resolution of the servo might be a problem at slow flows. | Measure start and end position with a caliper and compare with the volume recorded in the software. | ml | 5 | 0 | 5
2.3 | Check outputted volume after sine with amplitude 100 ml/s and frequency 6 rad/s for 3 periods. | Tests the backlash of the system for rapid changes. | Measure start and end position with a caliper and compare with the volume recorded in the software. | ml | 5 | 0 | 5
2.4 | Check outputted volume after sine with amplitude 10 ml/s and frequency 1 rad/s for 3 periods. | Tests the backlash of the system for slow changes. | Measure start and end position with a caliper and compare with the volume recorded in the software. | ml | 5 | 0 | 5
3. Mechanical tolerances
4. Software limits